Measurement and Modeling of Electron Cloud Trapping
in the CESR Storage Ring
Jim Crittenden
Cornell Laboratory for Accelerator-Based Sciences and Education
Low Emittance Rings 2014 Workshop
INFN-LNF, Frascati (RM), Italy
17 September 2014

---
Time-resolving shielded stripline
electron cloud detector
Collector
94.4 mm
7.1 mm

---
Discovery of electron trapping
June 2013
Shielded stripline signals
for 10- and 20-bunch trains
of 5.3 GeV positrons
in the 7 T/m quadrupole magnet
Shielded button electrode signals
recorded simultaneously
in a field-free region
of the CESR ring
Observation of Electron Trapping in a Positron Storage Ring,
M.G. Billing et al, arXiv 1309.2625
Shielded Button Electrodes for Time-resolved Measurements of Electron Cloud Buildup
JAC et al, Nucl. Instrum. Meth. A749, 42 (2014)
How do the first ten bunches know that ten more are coming?
Electron cloud which can contribute to the signal is trapped during the 2.3 ms interval between trains.
The change in slope indicates that six bunches suffice to clear this signal-producing trapped cloud.

---
Direct Measurement with Clearing Bunch
April 2014
Blue: 20-bunch train signal
Cyan: 20-bunch train with clearing bunch following at 900 ns
Magenta: Signal from a single bunch plotted to coincide with the clearing bunch for comparison
Cloud is present 900 ns after train and is accelerated into the detector by the clearing bunch.
The clearing bunch reduces the cloud available to make signal during the following train passage.

---
Shielded stripline signal modeling with
ECLOUD
Electron Cloud Buildup Characterization Using Shielded Pickup Measurements and Custom Modeling Code at CESRTA, JAC and J.P. Sikora, ECLOUD12, CERN 2013-002, p. 241, arXiv 1307.4013
* Originated at CERN in the late 1990's
   * Widespread application for LHC, KEK, RHIC, ILC …
   * Under active development at Cornell since 2008
   * Successful modeling of CESRTA tune shift measurements 2010
* Detailed photoelectron and SEY models developed 2010-2012
* Synrad3D photon absorption model included as input 2011
* Interactive modeling of time-resolved detector signals 2011 - 2013
   * Quadrupole stripline detector model implemented in 2013
Generation of photoelectrons
Production energy, angle
Azimuthal distribution (v.c. reflectivity)
Time-sliced cloud dynamics
Cloud space charge force
Beam kick
Magnetic fields
Secondary yield model
True secondaries (yields > 1!)
Rediffused secondaries (high energy)
Elastic reflection (dominates at low energy)
Model for a stripline detector in a quadrupole field
Acceptance vs incident angle, energy, B-field
Charge entering holes removed from cloud
Charge hitting wall creates secondaries
Snapshot of Modeled Trapped Cloud

---
Analytic form of the magnetic-field-dependent acceptance of the shielding holes
Numerical Modeling for CESRTA Measurements of Electron Cloud Buildup in a Quadrupole Magnet, JAC et al, TUPRI034, IPAC14
Examples
of
accepted
trajectories
B|| = 0.35 T
LH
No magnetic field
Courtesy of W. Hartung
Direct relationship between incident energy, angle qi and cyclotron radius Rc, transit time
Closed form obtained for arbitrary hole aspect ratio LH:RH and magnetic field magnitude (!)
Maximum accepted angle for zero field given by LH:RH = 3:1 (qi<20 degrees)
Acceptance for glancing incidence at high field and low energy

---
Shielded stripline signal modeling with
ECLOUD
Modeled signal during the third train passage
Signal charge is the incident charge weighted with the hole transmission factor
Voltage obtained from the amplifier input impedance of 50 W and +20 dB gain
A noise-reducing 13-MHz low-pass filter is applied to model and measurement, resulting in correlated error bars
Charge which does not enter the holes produces secondaries
0.641011 e+/bunch
1.001011 e+/bunch
1.161011 e+/bunch
1.321011 e+/bunch
The modeled signal is very sensitive to the details of the SEY parameters, such as the relative contributions of the true, rediffused and elastic secondaries, the shape of the true secondary dependence on incident energy, the energy distribution of the produced secondaries, etc.
The modeled signal is very sensitive to the details of the SEY parameters, such as the relative contributions of the true, rediffused and elastic secondary production processes, the shape of the true secondary yield dependence on incident energy, the energy distribution of the produced secondaries, etc.

---
Modeling the effect of the clearing bunch
Observation of Electron Trapping in a Positron Storage Ring, M.G. Billing et al, arXiv 1309.2625
The model tuned to reproduce the measured signals is used to infer the general characteristics of the electron cloud.
Here we consider the effect of the clearing bunch on the total cloud density in the beampipe.
Without the clearing bunch about 7% of the cloud is trapped.
The clearing bunch reduces the trapped cloud density by about a factor of four.

---
Modeling for the Positron Damping Ring of the International Linear Collider
SYNRAD3D Photon Propagation and Scattering and Simulations, G.F. Dugan and D. Sagan, ECLOUD12, CERN 2013-002, p. 117
Investigation into Electron Cloud Effects in the International Linear Collider Positron Damping Ring, JAC et al, Phys Rev ST Accel Beams 17, 031002 (2014)
Observation of Electron Trapping in a Positron Storage Ring, M.G. Billing et al, arXiv 1309.2625
The cloud density within 20 sbeam of the beam can reach 1.81012 e-/m3 in the quadrupole magnets in the wiggler region of the ILC positron damping ring.
Optics-distorting electric field gradients arising as the beam attracts cloud can reach values as high as 50 kV/m2.
Higher multipole fields should be taken into account as well.

---
Modeling for SuperKEK-B (Preliminary)
Sync rad photon rate in final focus quadrupole QC1RP
Design and Construction of the SuperKEKB Vacuum System, Y. Suetsugu et al, J. Vac. Sci. Technol. A30, 031602 (2012)
Study of Electron Cloud Effects in SuperKEKB, K. Ohmi and D. Zhou, proceedings of IPAC2014, Dresden, Germany, p. 1597
Assuming no photon scattering in vacuum chamber
Including Synrad3D photon scattering simulation
Impact position
of photons
Absorbed in
QC1RP
No photons
hit the v.c. wall
in QC1RP
directly
Number
of
reflections
Source position
of photons
Absorbed in
QC1RP
6-mm bunch length
Photon impact and
beam bunch
desynchronized
Detailed vacuum chamber profile courtesy of T. Ishibashi
Calculations up to now have not taken into account photon scattering inside the beampipe.
Ohmi & Zhou (IPAC14) found the tune shift to be dominated by electron cloud near, but not in, the IR region.
Using the LER lattice and the detailed wall profile, we find the rate of photons absorbed in the QC1RP final focus quadrupole magnet (66 T/m !) to be rather high: 1 photon/m/e+.

---
Electron Cloud Buildup Modeling for SuperKEK-B (Preliminary)
Electron-cloud-induced field gradients on the
beam axis in QC1RP reach values of 20 kV/m2.
The corresponding vertical tune shift is
DfY = c Y (dEY//dY) LQC1RP / ( 4  Ebeam Lring)
with Y =2500 m, LQC1RP=0.334 m, Ebeam=4e9 eV, and Lring=3016 m
DfY = 133 Hz
This value can be compared to the 52.7 kHz fractional vertical tune in SuperKEKB.
QC1RP is one of four final focus quadrupole magnets.
Multi-turn analysis to investigate effect of electron trapping remains to be done.
ECLOUD simulation of electron cloud buildup in the QC1RP final focus quadrupole for 4-ns spaced bunches of population 9.61010 e+ shows the density within 20 sbeam of the beam can reach 51012 e-/m3.

---
Conclusions and Next Steps
Summary
Electron trapping due to the magnetic mirror mechanism has been observed for the first time in a storage ring.
The measurements are well reproduced by modeling based on synchrotron-radiation-induced electron cloud buildup.
Applications of the model to the ILC positron damping ring and the low-energy SuperKEKB positron storage ring are likely to provide useful information on electron-cloud-related performance limitations.

---
Background
Magnetic mirroring is a well-known phenomenon (plasma physics)
Longitudinal and transverse momentum exchange as charges spiral along field lines
Adiabatic condition with conserved magnetic moment
Chaotic behavior in weak field regions
Escape zones and high-trapping zones (near, but not on, symmetry axis)
Coupled-bunch instability in KEK-B (L.Wang et al, Phys Rev E66, 036502 (2002))
Observed even with cloud-mitigating solenoids on
Multi-turn electron cloud trapping in quadrupole and sextupole magnets considered as source
Modeled in detail, but link to instability inconclusive
CESRTA “lead-bunch blowup” phenomenon
First bunch in 30-bunch train of positrons exhibits anomalous large vertical size
Reduced by adding a bunch near the end of the train or just before train
Motivated design and implementation of cloud detector in a quadrupole magnet
First measurement of electron trapping in a positron storage ring (M.G. Billing et al, arXiv:1309.2625)
Summer 2013 discovery by comparing signals from 10- and 20-bunch trains in a quadrupole magnet (7 T/m)
April 2014 studies of dependence on bunch population and spacing, and mitigation using clearing bunches
Extensive modeling to explain trapped cloud clearing during train passage, sensitivity to
secondary electron emission model, mitigating effect of clearing bunch or bunches,
time-dependence of cloud profile, nonlinear relationship of signal and cloud density
ILC positron damping ring study finds high cloud densities in wiggler-region quadrupoles (10 T/m)
(JAC et al, Phys Rev ST Accel Beams 17, 031002 (2014))
Heat load from trapped cloud in LHC final-focus quadrupoles identified as limit on future LHC operation
(G. Iadarola, G. Rumolo and C. Zannini, 2014, in preparation)
Begin study of electron cloud in SuperKEK-B final focus quadrupole magnets (70 T/m !)
(JAC, T. Ishibashi, in progress)

---
Time-resolving shielded stripline
electron cloud detector
 Stripline frequency response measured with a spectrum analyzer
- Cutoff frequency is 13 MHz (RC=12.2 ns)
- High-pass feature results in sensitivity to beam-induced noise
Unfiltered stripline signal
12.2 ns RC filter

---
Observation of Cloud Buildup Resonant with Bunch Spacing
Signals from 20-bunch trains with spacings of 14, 16, 20, 24, and 28 ns are shown. The average bunch populations are 1.3x1011 positrons. The signal generally decreases as the bunch spacing is increased. However, the signal increases when the bunch spacing is increased from 14 ns to 16 ns, indicating a resonance in the production of cloud electrons.
Be careful when choosing bunch spacing !

---
Shielded stripline signal modeling with
ECLOUD
 δSEY (E=Epeak) = 1.53
 δSEY (E=Epeak) = 1.63
The modeled signal is very sensitive to the details of the SEY parameters, such as the relative contributions of the true, rediffused and elastic secondaries, the shape of the true secondary dependence on incident energy, the energy distribution of the produced secondaries, etc.
0.641011 e+/bunch
1.001011 e+/bunch
1.161011 e+/bunch
1.321011 e+/bunch
0.641011 e+/bunch
1.001011 e+/bunch
1.161011 e+/bunch
1.321011 e+/bunch

---
Trapped signal electrons produced near the end of a preceding train
Since six bunches suffice to clear the trapped cloud which can produce signal, such electrons can only be produced during the trailing six bunches of a train
4.01 mA/bunch
Trapped signal electrons produced between trains
Trapping for signal electrons occurs more effectively at lower bunch current
8.26 mA/bunch

---
Striking divergence of the beam-pipe-averaged cloud density from 7.26 to 8.26 mA/bunch
The cloud density increases nearly linearly with bunch current from 4 to 7 mA, then shows an increase of a factor of five from the first train to the second, and an additional factor of two from the second to the third train passage.
Ratios: 1 : 1.6 : 2.4 : > 36 (!)
1) The assumption of asymptotic behavior after only two turns is now questionable.
2) Why doesn't the modeled signal show this behavior?
4.01 mA
6.26 mA
1.50e11
2.40e11
2.05e9
2.98e9
7.26 mA
8.26 mA
2.85e11
1.10e12
4.95e9
3.95e10
NB: The peak densities differ from the axis values owing to bin averaging.

---
Where is the cloud hiding?
7.26 mA
8.26 mA
Q48W Detector Position
Trapped cloud at the end of the second turn
7.26 mA: 2.8e7 e
8.26 mA: 2.2e8 e
Cloud immediately following the 60th filled bunch
7.26 mA: 1.3e9 e
8.26 mA: 9.4e9 e

---
Where do the signal electrons enter the detector?
4.01 mA
6.26 mA
7.26 mA
8.26 mA
Potential signal electrons are missing the detector !
We should consider building a detector with more acceptance and transverse segmentation.

---
Modeling for SuperKEK-B
(Preliminary)
Design and Construction of the SuperKEKB Vacuum System, Y. Suetsugu et al, J. Vac. Sci. Technol. A30, 031602 (2012)
No photon scattering in vacuum chamber
Calculations up to now have not taken into account photon scattering inside the beampipe.
Ohmi (IPAC14) found the tune shift to be dominated by electron cloud near, but not in, the IR region.
Using the LER lattice and the detailed wall profile provided by T. Ishibashi/KEK, we find the rate of photons absorbed in the QC1RP final focus quadrupole magnet (66 T/m) to be rather high: 1 photon/m/e+.

---
SYNRAD3D Modeling for SuperKEK-B
(Preliminary)
Synrad3D finds scattered photons absorbed in the QC1RP final focus quadrupole which
originated in dipole magnets up to 36 m upstream and having undergone 1-20 scatters off the inner wall of the vacuum chamber.
Quite a few of the photon trajectories differ from the beam trajectory by more than the 6 mm bunch length.